## Quotation
# | Suggested Description to order | Unit | Quantities Requested | Unit price in $ (Including Customs & Vat & All Other Taxes & Hidden Costs) | Total | Delivery Time (In Weeks)
A | N95 Particulate Respirator Mask (Large) | Piece | 6000
A | N95 Particulate respirator Mask (Medium) |  | 6000
B | 3 Ply Surgical Mask, | Box of 50 | 8000
C | Disposable Surgical Gown | Piece | 180000
D | Goggles / Eye Protection | Piece | 75000
E | Suits Coverall-Large-L | Piece | 9500
E | Suits Coverall-Xtra Large-XL |  | 10000
E | Suits Coverall-XXL |  | 8500
E | Suits Coverall-XXXL |  | 8000
F1 | Gloves Latex-Medium | Box of 100 | 2300
F2 | Gloves Latex-Large | Box of 100 | 2300
O | Cover Boots | Pair | 7000
P | Gloves Nitrile-Large | Box of 100 | 1200

## Specifications
# | Description | Specifications | Required with the quote | Unit
A | N95: Particulate respirator compliant with NIOSH N95 or EN 149 FFFP2 and fluid resistant (1/2 medium, 1/2 Large) | Material: non-woven filter layer, fluid resistant, Filtration level: > 95 % for particles from 0.1 to 0.3 micron,Air permeability: > 2 mm H2O,Shape of the mask: duckbill, folded (horizontal) width-wise Without valve, Respirator mask fits all face shapes, without inspiration/expiration air-leakage Respirator mask conforms to standards: NIOSH N95 or EN 149: 2001/FFP2, or EN 149:2001/FFP3 The disposable respirator should be marked with the manufacturer's name, the part number (P/N), the protection provided by the filter (e.g. N-95), and "NIOSH." or EN 149: 2001/FFP2 Or EN 149:2001/FFP3 and CE certification number. | Manufacturer Name, Certification for standards used, Specification Data Sheet with Laboratory certification Test and a Photo of the product. | Piece
B | 3 ply surgical face mask, compliant with EN 14683 type IIR or ASTM F2100 level2 or level 3 or equivalent | Material: non-woven with outer layer impervious liquid splash resistant material Mask comprises 3-4 non-woven layers, and has clearly identifiable inner and outer surfaces Size: Approx. 17x10cm (l x w) Single-use disposable Compliant with EN 14683 type IIR or ASTM F2100 level2 or level 3. | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet and a Photo of the product. | Box of 50
C | Disposable gown: single use, long sleeves, fiber made non-woven, thumb loop, tape tab for neck closure, water and liquids proof, compliant with the EN 13795 high performance level, or AAMI level 3 performance or equivalent | Surgical gown, Single Use with long sleeves, and a waist tie that binds at the side Fabric, non-woven material, f.e. SMS, SMMS, etc. Outer layer liquid penetration resistant. Sleeves finished with double layer cuff, cotton or synthetic, stretchy interlocked jersey, length: 4 - 8 cm, Single use, disposable,Size: L Conform requirements of: (a) EU Medical Devices Directive 93/42, Class I or (b) US FDA 510(k) clearance, Class II or (c) Equivalent internationally recognised marketing clearance In specific, compliant with: (a) EN 13795 standard (in the standard or high performance category) or (b) ANSI/AAMI PB70:2012 standard (in the AAMI Level 3 or AAMI Level 4 category). | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet. And a photo of the product | Piece
D | Goggles / eye protection compliant with EU standard directive 86/686/EEC, EN 166/2002 or ANSI/ISEA Z87.1-2010 or equivalent | Material, frame: translucent soft PVC. -Indirect ventilation channels (preventing penetration of splashes), one through each side of the frame.,Diameter of each vent, at least: 20mm.,Wrap-around adjustable elasticated headband, integrated with Material, lens part: polycarbonate, thickness 1 mm.,Lens is anti-fog treated/coated. Compliant with EU standard directive 86/686/EEC, EN 166/2002 or ANSI/ISEA Z87.1-2010 for chemical splash and dust protection | Manufacturer Name, Certification for standards used, Specification Data Sheet. And a photo of the product | Piece
E | Disposable coverall: compliant with EN 943-1:2002 such TYVEK or equivalent (1/3L, 1/3XL, 1/3XXL) | European Directive 89/686/EEC on personal protective equipment Category III: Chemical protective coverall, Type 6 comply with EN 13034:2005+A1 Elasticated hood around face,Elasticated cuffs and ankles, Elasticated sleeves,Preferably bound seams,Zipper. through seams,Stitched-in neck label indicating the type and performance of the suit against the below mentioned standards. Color: White Material: Lightweight, do not contain rubber/ latex. Antistatic treated on both sides. Fabric is Infective agent tested against viral penetration at minimum 1.75kPa (class 2) Non-sterile,Single Use, disposable Barrier to infective agent standards: EN 14126:2003 certified passing infectious agent test according to ISO 16604:2004 (Resistance to penetration by blood‐borne pathogens using bacteriophage Phi‐X 174) standard at minimum exposure pressure of 1.75kPa (class 2) (or equivalent international standard) Performance requirement ISO standards: EN 340 – General requirements,EN ISO 17491-4: 2008 (supersedes: EN 468 mod. – Spray aerosol test),EN 14325 – Test methods & performance classification,EN ISO 3758 (ISO 3758) – Textile care symbols,EN ISO 13935-2 – Seam strength | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet. And a photo of the product | Piece
F1 | Latex gloves: powder free / hypoallergenic / tear resistant/ sterile/single use/ambidextrous/ISO 9001 - Large size | Examination gloves, ambidextrous shape: fits either hand.Material: Natural latex, Powder-free,Total length: approximately 230mm,Palm width: approximately 110 + or - 10mm.Single-use,Non sterile,CE marked / Conform to 93/42/EEC Class I, EN-455 or ISO 9001 | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet. | Box of 100
F2 | Latex gloves: powder free / hypoallergenic / tear resistant/ sterile/single use/ambidextrous/ISO 9001 - Medium size |  | Manufacturer Name, Origin Certification for standards used, Specification Data Sheet. | Box of 100
O | Cover shoes for tyvek coverall or equivalent: closed / impermeable / length to below knee / washable / disinfected if reusable | Closed / impermeable / length to below knee / washable / disinfected if reusable/ Conform to 99/92/CE , EN 1149 | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet and Photo of the product | Pair
P | Nitrile Gloves, Large Size | Nitrile, Powder free, non-sterile, Single-use disposable, Fits either hand, Thickness at fingertips of ≥ 120um, Box of 100 gloves marked Conform to 89/686/EEC Category III, EN 374 | Manufacturer Name, Origin, Certification for standards used, Specification Data Sheet. | Box of 100